## Slide 1: BITNG PROJECT UPDATE
Carl Demolder
Date 10/22/2020

## Slide 2: Outline
Progress to date
Shriner’s project
Schedule
  Gantt Chart update
Path forward

## Slide 3: PROGRESS TO DATE

## Slide 4: Progress from last week
Pediatric wearable
  Literature review [INITIAL TABLES]
  IRB proposal [INTIAL DRAFT]

## Slide 5: SHRINER’S PROJECT

## Slide 6: PCB design
PCB: Arrived
Stencil: Arrived
Digikey: Arrived
Assembly: Complete
Hardware Debug: TO DO
Programming Check: TO DO

## Slide 7: LP ECG PROJECT

## Slide 8: PCB design
PCB: Ordered
Stencil: Arrived
Digikey: Shipped

## Slide 9: CAPACITANCE SENSOR

## Slide 10: Capacitance sensor
PDMA: Finished
PI: Finished
AG: Finished
PDMS: Finished
ASSEMBLY: TO DO
ENCAPSULATION: TO DO

## Slide 11: IRB PROPOSAL

## Slide 12: IRB proposal
Initial Draft:
  IRB Protocol: 80% Complete
  Final Questionnaire
  Recruitment Flyer
  Consent Form
  COI Management Plan
  Device Statement
  Device Biocompatibility
  Device Biocompatibility Reference Article
Timeline:
  Initial Draft sent to Dr. Yeo on Tuesday (10/13)

## Slide 13: PATH FORWARD

## Slide 14: Path forward (10/19/20 – 10/26/20)
Pediatrics Wearable:
  Literature review
    Add more sources to tables
PCB Assembly:
  Wearable Sensor Glove
  LP ECG
Sensor Manufacturing:
  Capacitance sensors (10X)

## Slide 15: SCHEDULE

## Slide 16: Schedule Gantt chart
Task | 10/11-10/18 | 10/18-10/25 | 10/25-11/01 | 11/01-11/08 | 11/08-11/15 | 11/08-11/15 | 11/16-11/23
LOW POWER ECG |  |  |  |  |  |  | 
-PCB ASSEMBLY |  |  |  |  |  |  | 
-HARDWARE DEBUGGING |  |  |  |  |  |  | 
-FIRMWARE DEVELOPMENT |  |  |  |  |  |  | 
NEUROMOTOR PEDIATRIC WEARABLE |  |  |  |  |  |  | 
-LITERATURE REVIEW |  |  |  |  |  |  | 
-PCB ASSEMBLY |  |  |  |  |  |  | 
-HARDWARE DEBUGGING |  |  |  |  |  |  | 
-FIRMWARE DEVELOPMENT |  |  |  |  |  |  | 
-PRESSURE SENSOR MANUFACTURING |  |  |  |  |  |  | 
-IRB PROPOSAL |  |  |  |  |  |  | 
FINSHED
TO-DO
LEGEND

## Slide 17: APPENDIX

## Slide 18: Capacitor sensors
Data collected from Robby
Max capacitance change
  20% ~ 10 pF
Increasing pressure

## Slide 19: Capacitor sensors
Data collected from Robby
Sensitivity: 0.133 pF/kPa
Pressure threshold > 0.5 kPa
example data
capacitance change
percent change

## Slide 20: Semmes Weinstein scale
4-180 grams of target force

## Slide 21: Capacitance sensor design
Capacitor area (prior art): 5mm x 5mm
Initial capacitance: 52 pF
Sensitivity: 0.133 pF/ kPa
 | Low threshold | High Threshold
Target force (g/N) | 4 / 0.0392 | 180 / 1.764
Target pressure (kPa) | 1.568 | 70.56
Capacitance change (pF/%) | 0.2085 / 0.4 | 9.384 / 18
CDC units | 910 | 40,999

## Slide 22: Block diagram-motherboard
SENSORY
SENSING
WEARABLE
Wired Digital Communication
NRF52832
FEEDBACK
ARMBAND
MOTHERBOARD
LIPO
BATTERY
BMS
INDUCTIVE
WIRELESS
CHARGING
SENSORY
SENSING CIRCUIT

## Slide 23: Block diagram-sensory sensing
SENSORY
SENSING
PRESSURE
SENSORS [3X]
CAPACITIVE
PRESSURE
SENSORS [3X]
CAPACTIVE
SENSING
CIRCUIT [3X]
FINGER STRAIN
SENSOR [3X]
RESISTANCE
STRAIN
SENSORS [4X]
RESISTIVE
SENSING
CIRCUIT [7X]
GRASP STRAIN
SENSOR [1X]
WIRED DIGITAL COMMUNICATION [4 LINES]
NRF52832
TEMPERATURE
SENSOR [3X]
RESISTANCE
TEMPERATURE
SENSORS [3X]
SENSOR GLOVE
MOTHERBOARD

## Slide 24: PCB design
~60mm x ~40mm

## Slide 25: Resistive sensing circuit
Up to 8 resistance sensors

## Slide 26: Capacitive sensing circuit
Up to 4 Capacitance sensors**

## Slide 27: Power circuit

## Slide 28: MCU circuit

## Slide 29: Wireless charging: Inductive
Hardware
  Old wireless charging coil:
    TDK 45-15F5-G (48.2mm x 32.2mm x 1.0mm)
  Current inductive coil:
    WR222230-26M8-G (22 mm x 22mm x 0.87 mm)
TRANSMITTER
RECEIVER COIL
RECEIVER
CIRCUITRY
FOOTPRINT:
15mm x 15mm

## Slide 30: Wireless charging: RF
RF Harvesting
  Battery and battery-less electronics
  7 mA continuous @ 0.5 meter
Companies contacted for development kit:
  Humavox
    Still waiting for a response
  Energous
    In conversation to receive 10 units of their receivers
  Powercast
    Digikey [PURCHASED]
      Transmitter (TX91503 $99.00)
        Battery-less receiver module (P2110B $48.33)
        Battery charging receiver module (P1110B $32.00)
        Development kit (P1110-EVAL-PS $420.00)

## Slide 31: HOUSEKEEPING

## Slide 32: Housekeeping
Electronics inventory:
  Google spreadsheet listing what is in stock at the lab
  Resistors, capacitors, inductors, ICs, etc…
  Reduce wait time to purchase parts

## Slide 33: Electronics development: schematic
Schematic generation
  nRF52 with supporting circuitry (2 hours)
  BLE antenna and impedance matching circuit (2 hours)
  Wireless charging (5 hours)
  Lithium battery charging (2 hours)
  UI interaction (1 hour)
  USB charging (2 hours)
  Biosensors: temperature, IMU, ECG (4 hours)
  Test points, connectors, jumpers (2 hours)
Estimated time: 20 hours
Actual time spent: 22 hours
*Not in scope for current phase of development.

## Slide 34: Progress to date: 6/28/2020

## Slide 35: Progress to date: 6/28/2020

## Slide 36: Progress to date: 6/28/2020

## Slide 37: Progress to date: 6/28/2020

## Slide 38: Electronics development: PCB design
PCB design
  Major component layout (2 hours)
  BLE antenna and impedance matching circuit (3 hours)
  Wireless charging (2 hours)
  Lithium battery charging (1 hours)
  Overall board construction(1 hour)
  USB charging (2 hours)
  Biosensors: temperature, IMU, ECG (3 hours)
  Test points, connectors, jumpers (1 hours)
  Gerber file generation (2 hours)
  BOM generation (2 hours)
Estimated time: 16 hours
Actual time spent: 20 hours

## Slide 39: Progress to date: 6/28/2020

## Slide 40: Progress to date: 6/28/2020

## Slide 41: Progress to date: 6/28/2020

## Slide 42: PCB manufacturing specifications
Dimensions: 99.568 x 76.454 mm
Layers: 2
Minimum trace/space: 6 mil
Minimum hole size: 12 mil
Thickness: 1.6mm
Solder mask: Green
Surface finish: HASL with lead
Copper weight: 1 oz
Material type: FR4

## Slide 43: PCA quote
PCB Manufacturing:
  Price: $20
  Manufacturer: PCBMinions
  Lead time: 4 days
BOM:
  Price: ~$289.63
  Board Quantity: 3
  Supplier: Digikey/Mouser
  Lead time: 1-5 days
Stencil:
  Price: ~46.25
  Supplier: OSH Stencil
  Lead time: 5 days

## Slide 44: Firmware development
Firmware
  Project and BLE stack (5 hours)
  Sensor drivers
    I2C driver (4 hours)
    Temperature driver (2 hours)
    ECG driver (6 hours)
    IMU driver (3 hours)
  USB & RTT data logging (5 hours)
    RTT data logging (2 hours)
    USB data logging (3 hours)
  DAQ slave protocol (1 hour)

## Slide 45: PROJECT MANAGEMENT

## Slide 46: Introduction
Purpose:
  Design a development board to test the functionality of different ‘plug-in-play’ sensors, wireless charging, and BLE communication. In the future, these sensors will be used in different combinations to build wearable devices.

## Slide 47: Project management
Project status updates will be emailed to YSK and Dr. Yeo on a weekly basis by Monday morning (ET)
A rolling PowerPoint presentation will be used to document results and act as a final report
  Each progress update will be timestamped, and the revision number will be incremented
Primary form of communication: email
  Zoom/phone calls can not be conducted from 12 PM – 8 PM ET, Monday – Friday (due to my work schedule)
Project documents and resource will be available through my Github repository
Time commitment: average 10 hours/week

## Slide 48: Project requirements
Requirements:
  Design a development board PCB to demonstrate capabilities to connect to different bio-sensors
  Development board must include the following functionality:
    Wireless and USB charging
    Battery management
    USB data communication and logging
    Use the nRF52832 MCU/SoC
    BLE connectivity and BLE GATT database
    Temperature sensing
    Motion sensing
    ECG
    Capability to add more sensors in the future

## Slide 49: Schedule overview
Electronics development
  Date range: 6/1–7/31
  Estimated hours to produce hardware: 49 hours
  Estimated hours to test/debug/qualify hardware: TBD
Firmware development
  Date range: 6/28-7/31
  Estimated hours to produce firmware: 33 hours
  Estimated hours to test/debug/qualify firmware: TBD
Any questions about the scope?
Let me know!

## Slide 50: PROJECT PROPOSAL

## Slide 51: Firmware development
Firmware
  Project and BLE stack (5 hours)
  Sensor drivers
    I2C driver (4 hours)
    Temperature driver (2 hours)
    ECG driver (6 hours)
    IMU driver (3 hours)
  USB & RTT data logging (5 hours)
Future development*
  BLE GATT database and data transfer (25 hours)*
  Sensors: fuel gauge, H.R./Sp02 (8 hours)*
  Data storage: EEPROM, Flash (8 hours)*
*Not in scope for current phase of development.

## Slide 52: Hardware testing
Firmware testing
  Firmware sensor functionality testing (8 hours)
Hardware testing
  Overall hardware testing (5 hours)
  Hardware sensor functionality testing (8 hours)
Future development*
  Power consumption*
  BLE range and connection testing*
*Not in scope for current phase of development.

## Slide 53: Block diagram
PYTHON DATA LOGGER
BLE TX/RX
IMU
TEMPERATURE
MCU/SoC
SERIAL TO USB CONVERTER
ECG
HR/Sp02**
PHERIPHERAL SENSORS
BATTERY GAUGE**
SERIAL
WIRELESS
CHARGING COIL
WIRELESS CHARGING CIRCUITRY
BMS
LIPO BATTERY
SERIAL
USB
POWER
POWER
POWER
POWER
RF
**Not in scope for current phase of development.
ANDROID PHONE**
BLE
OUT OF SCOPE

## Slide 54: PROJECT PROPOSAL: POWER CONSUMPTION TESTING
OUT OF SCOPE

## Slide 55: Power consumption testing [TBD]
Objective: Measure the power consumption of various sensors during operation, idle, and sleep mode
Log battery current and voltage at a defined sampling rate
Produce a power consumption model to determine battery life
OUT OF SCOPE

## Slide 56: Power consumption testing [TBD]
Nordic power profiler
  Nano-amp resolution, fast sampling rate
  See spikes of BLE transmission
OUT OF SCOPE

## Slide 57: Schedule flowchart: hardware
PROJECT
PROPOSAL
SCHEMATIC
GENERATION
PCB COMPONENT
LAYOUT
PCB ROUTING
& MFG. FILE GENERATION
FULL TURNKEY
PCB MFG.
HARDWARE
TESTING*
POWER CONSUMPTION TESTING*2
SENSOR FUNCTIONALITY TESTING*1
6/1–6/7; ~5 Hours
6/7–6/14; ~20 Hours
6/14–6/17; ~3 Hours
7/12–7/19; 5 Hours
7/19–7/31; 8 Hours
6/28–7/12; ~4 Hours
6/17–6/28; ~12 Hours
8/23–9/6; TBD
*Physical hardware needed to debug and perform testing.
1Hardware functionality test of the sensors will be performed. Performance and accuracy will not be qualified. 2Not in scope for current phase of development.
BLE RANGE & CONNECTION
TESTING*2
8/23–9/06; TBD
OUT OF SCOPE

## Slide 58: Schedule flowchart: firmware
PROJECT
PROPOSAL
PROJECT STACK
ORGANIZATION
SENSOR
DRIVERS
SERIAL & RTT DATA LOGGING W/ PYTHON
SENSOR FUNCTIONALITY TESTING*1
BLE GATT
DATABASE2
POWER
CONSUMPTION
TESTING*1,2
7/1–6/7; ~5 Hours
6/28–7/2; ~5 Hours
7/2–7/12; ~15 Hours
8/16–8/23; ~25 Hours
8/23–9/6; TBD
7/19–7/31; ~8 Hours
7/5–7/12; ~5 Hours
*Physical hardware needed to debug and perform testing.
1Firmware functionality test of the sensors will be performed. Performance and accuracy will not be qualified.
2Not in scope for current phase of development.
OUT OF SCOPE

## Slide 59: Component selection: MCU/BLE
MCU/BLE: nRF52832
  6 pin, 2.00 mm x 2.00 mm, WSON-6
  ±0.1°C from -20°C to +50°C
  2.4 GHz BLE
  Power = 4 dBm
  Sensitivity = -96 dBm
  RAM = 64 kB, flash = 512 kB
  Serial interfaces: I2C, UART, SPI
  Supply voltage: 1.7V ~ 3.6V
  $5.46 [X1], 1490-1067-1-ND, LINK
nRF52832

## Slide 60: Component selection: clocks
HF clock: US3200005Z
  4 pin, 1.60 mm x 1.20 mm, 4-SMD
  32.00 MHz, 8 pF, ±10 ppm
  $1.45 [X1], US3200005ZCT-ND, LINK
LF clock: ECX-1210
  1210 package, 1.20 mm x 1.00 mm, 2-SMD
  32.7680 kHz, 9 pF, ±20 ppm
  $1.52 [X1], XC2283CT-ND, LINK
ECX-1210
US3200005Z

## Slide 61: Component selection: USB/Serial
USB connector: 62910416121
  4 pin, 16.00 mm x 10.62 mm, SMT type B
  USB TYPE-B, USB 2.0, right angle
  $3.83 [X1], 732-3152-1-ND, LINK
USB/I2C converter: FTDI FT201XQ-R
  16 pin, 4.0 mm x 4.0 mm, 16-QFN
  USB to I2C converter
  3.4 Mbit/s
  Low power consumption (8 mA)
  $2.04 [X1], 768-1131-1-ND, LINK
FT201XQ-R
62910416121

## Slide 62: Component selection: Bio-sensors
Temperature: TI TMP117
  6 pin, 2.00 mm x 2.00 mm, WSON-6
  ±0.1°C from -20°C to +50°C
  16-bit resolution: 0.0078 °C, I2C
  $8.02 [X1], 296-51834-1-ND, LINK
IMU: TI BQ27621-G1
  14 pin, 2.5 mm x 3.0 mm, 14-LGA
  I2C communication
  9-axis, 16-bit resolution
  Full operation mode = 925 uA
  $4.98 [X1], 828-1057-1-ND, LINK
BMI160
TMP117

## Slide 63: Component selection: Bio-sensors
H.R. & Sp02: MAX30102
  14 pin, 5.6 mm x 3.3 mm, 14-OESIP
  Pulse oximetry and heart-rate
  Temperature compensation, I2C
  $9.51 [X1], MAX30102EFD+-ND, LINK
ECG: MAX30003
  28 pin, 5.0 mm x 5.0 mm, 28-TQFN
  SPI communication
  15.5-bit resolution
  Ultra-low power, single-channel integrated
  $8.21 [X1], MAX30003CTI+-ND, LINK
MAX30003
MAX30102
OUT OF SCOPE

## Slide 64: Component selection: wireless charging
Wireless charging IC: TI BQ51003
  28 pin, 1.90 mm x 3.00 mm
  QI v1.2 communication protocol
  93% AC-DC efficiency
  $3.00 [X1], 296-39952-1-ND, LINK
LiPo battery charger IC: TI BQ25100
  6 pin, 1.6 mm x 0.90 mm, 4.2V 1 cell
  1% Charge voltage accuracy
  10 mA to 250 mA charging current
  $2.24 [X1], 296-38373-1-ND, LINK
BQ25100
BQ51003

## Slide 65: Component selection: power circuit
Reed Switch: ABLIC S-5712A
  4 pin, 1.90 mm x 3.00 mm, SNT-4A
  Non-contact ‘power switch’
  High accuracy hall effect switch IC
  $1.07 [X1], 1662-1617-1-ND, LINK
Linear Regulator: Analog Devices ADP172
  4 pin, 0.95 mm x 1.03 mm, 4-WLCSP
  0.1V dropout voltage
  300 mA max current
  $1.09 [X1], ADP172ACBZ-3.0-R7CT-ND, LINK
ADP172
S-5712A

## Slide 66: Component selection: power circuit
Charging coil: Wurth 760308101219
  1 coil, 1 layer, D=15 mm, H=0.6 mm
  11.8 uH, 750 mΩ, t = ±10%
  $7.43 [X1], 732-9675-ND, LINK
Fuel gauge: TI BQ27621-G1
  9 pin, 1.62 mm x 1.58 mm, 9-DSBGA, 0.5 mm pitch
  I2C communication
  Dynamic voltage correlation algorithm
  Battery capacity, state-of-charge, voltage
  $2.21 [X1], 296-37460-1-ND, LINK
BQ27621-G1
760308101219
OUT OF SCOPE

## Slide 67: Component selection: TBD
Components still need to be selected:
  BLE antenna
  BLE impedance matching circuit/filter
  DCDC circuit

## Slide 68: Possible functionality
Possible functionality:
  Local data storage? (SD Card, EEPROM, etc…)
  Type of ECG connector? (Digikey part number…)
  NFC?
  Any other ideas?