## Slide 1: BITNG PROJECT UPDATE
Carl Demolder
Date 12/07/2020

## Slide 2: Outline
Progress to date
LP ECG
Schedule
  Gantt Chart update
Path forward

## Slide 3: PROGRESS TO DATE

## Slide 4: Progress from last week
LP ECG
  Firmware development
IRB Proposal
  Modification

## Slide 5: SHRINER’S PROJECT

## Slide 6: PCB design
PCB: Arrived
Stencil: Arrived
Digikey: Arrived
Assembly: Complete
Hardware Debug: Completed
Programming Check: TO DO

## Slide 7: LP ECG PROJECT

## Slide 8: PCB design
PCB: Arrived
Stencil: Arrived
Digikey: Arrived
Assembly: Completed
Hardware Debug: Completed
Programming Check: TO DO

## Slide 9: Firmware
ECG
  MAX30003
    Drivers
  Optimization
FRAM
  CY15B108QI
    Drivers
  Optimization
NRF52
  BLE
  Power Optimization
  RTC
Debugging

## Slide 10: GATT Server [ECG Patch]
Service | Characteristic | Data Type | Read/Write | Notification
Configuration | Settings | uint8_t [10] | Write/Read | NO
 | Response | uint8_t [2] | Read | Yes
 | CRC | uint8_t [2] | Read | Yes
Temperature | Temperature | uint8_t [250] | Read | Yes
ECG | ECG | uint8_t [250] | Read | Yes

## Slide 11: ECG Patch - Wakeup
Wave magnet over patch
Trigger Hall Effect Sensor
Enable LDO
VCC Supplied to MCU
Peripheral Initialization
Log Driver
NRFX Clock Driver
I/O Pins
Enable
LDO Drivers
Enable VCC LDO
Magnet Removed
Power Manager
<1 second
LED Driver
LED Blink
Init/Start
HF CLOCK
Init/Start
LF CLOCK
Init/Start
TWIM Driver
Init
TMP117 & RTC
Init/Start
SPIM Driver
Init
MAX30003
Init
CY15B108QI
CHECK IF SERIAL DEBUGGER ATTACHED
Serial Handler
Bluetooth
Handler
YES
NO

## Slide 12: Bluetooth Handler
Init BLE Stack
Init GAP Parameters
Init GATT
Init BLE Advertising
Set BLE Advertising Power
Init Connection Parameters
Init Peer Manager
Enable DC/DC Converter
Start Advertising
Shutdown
BLE Connection Established
LED Blink
GATT Client Discovers Services
Subscribe to characteristic notifications
GATT Client changes “Settings” Characteristic
GATT Server changes “Response” Characteristic to send request status and sends notification to GATT Client
GATT Server changes “Response” Characteristic to report error performing requested operation
Check connection before timeout
YES
NO
Change setting or perform action
FAILURE
SUCCESS
GATT Server writes “Response” Characteristic to “Initialization” Command & Sends Notification

## Slide 13: Data Recording Initialization
GATT Client Sends Command to Read Firmware and Hardware Version
GATT Client sends command to change sampling rate
GATT Client Sends Command to Shutdown
Disable/Uninit Unused Peripheral Drivers
Start MAX30003 Data Recording
LED Blink
Start RTC for TMP117
GATT Server Disconnects BLE
Enter Low Power Mode
Check if HW & FW Versions are Up to Date
NO
YES
Check Temperature and ECG Sampling Rate
NO
Check Recording Session Duration
YES
GATT Client Sends Command to Change Recording Session Duration
NO
GATT Client Sends Command to Start Recording Session
GATT Server Sends Confirmation Through “Response” Characteristic and Sends Notification
Wait for Interrupt

## Slide 14: Data Interrupt
Interrupt Occurred
Init/Start TWIM Driver
Read Temperature Value from TMP117
Write Data to Registers
Wakeup CY15B108QI
Read FIFO Memory
Wakeup CY15B108QI
TMP117 RTC or MAX30003?
TMP117
MAX30003
Reset FIFO Register
Init/Start SPIM Driver
Read Status Register
Write Data to Registers
Check FIFO Overflow
YES
NO
Deep Sleep CY15B108QI
Disable/Uninit Unused Peripheral Drivers
Enter Low Power Mode
Deep Sleep CY15B108QI
Disable/Uninit Unused Peripheral Drivers

## Slide 15: Recording Session Timeout
Recording Session Timeout
GATT Client Reads “ECG” Characteristic
Read Data off CY15B108QI
Wakeup CY15B108QI
Init/Start SPIM Driver
GATT Client reads “Temperature” Characteristic
BLE Connection
Shutdown
Start Advertising
GATT Server Writes Data to “ECG” Characteristic & Sends Notification to GATT Client
Check Data Available to Transfer
YES
NO
Disable/Uninit Unused Peripheral Drivers
Enter Low Power Mode
NO
Deep Sleep CY15B108QI
GATT Server Writes Data to “Temperature” Characteristic & Sends Notification to GATT Client
ECG Data Available for Transfer
Temperature Data Available for Transfer
YES
YES
NO
NO
GATT Server Writes “Response” Characteristic To “Data Ready” Command & Sends Notification

## Slide 16: STRAIN SENSOR

## Slide 17: Strain Sensor: Round #1
PDMA: Finished
PI: Finished
AG NW: Finished
PDMS: Finished
ASSEMBLY: FAILED
ENCAPSULATION: FAILED
Characterization: FAILED
Broken pads
Not enough nano-wires

## Slide 18: Strain Sensor: Round #2
PDMA: Finished
PI: Finished
AG NW: Finished
Conductivity Test: FAILED
PDMS: TO-DO
ASSEMBLY: TO-DO
ENCAPSULATION: TO-DO
Characterization: TO-DO

## Slide 19: PATH FORWARD

## Slide 20: Path forward (12/07/20 – 12/14/20)
LP ECG:
  Firmware development
  Flexible PCB design
Shriner’s Project:
  Sensor manufacturing:
    Strain sensors (2X)
      Trial #3

## Slide 21: SCHEDULE

## Slide 22: Schedule Gantt chart
Task | 11/23-11/30 | 11/30-12/07 | 12/07-12/14 | 12/14-12/21 | 12/21-12/28 | 12/28-1/4 | 1/4-1/11
LOW POWER ECG |  |  |  |  |  |  | 
-PCB ASSEMBLY |  |  |  |  |  |  | 
-HARDWARE DEBUGGING |  |  |  |  |  |  | 
-FIRMWARE DEVELOPMENT |  |  |  |  |  |  | 
-PCB DESIGN |  |  |  |  |  |  | 
NEUROMOTOR PEDIATRIC WEARABLE |  |  |  |  |  |  | 
-LITERATURE REVIEW |  |  |  |  |  |  | 
-PCB ASSEMBLY |  |  |  |  |  |  | 
-HARDWARE DEBUGGING |  |  |  |  |  |  | 
-FIRMWARE DEVELOPMENT |  |  |  |  |  |  | 
-PRESSURE SENSOR MANUFACTURING |  |  |  |  |  |  | 
-STRAIN SENSOR MANUFACTURING |  |  |  |  |  |  | 
FINSHED
TO-DO
LEGEND

## Slide 23: APPENDIX

## Slide 24: CAPACITANCE SENSOR

## Slide 25: Capacitance sensor
PDMA: Finished
PI: Finished
AG: Finished
PDMS: Finished
ASSEMBLY: Finished
ENCAPSULATION: Finished
Characterization: TO-DO

## Slide 26: Capacitance/Force characterization
Mark-10 Force Gauge
  Automated testing characterization

## Slide 27: Capacitor sensors
Data collected from Robby
Max capacitance change
  20% ~ 10 pF
Increasing pressure

## Slide 28: Capacitor sensors
Data collected from Robby
Sensitivity: 0.133 pF/kPa
Pressure threshold > 0.5 kPa
example data
capacitance change
percent change

## Slide 29: Capacitance sensor design
Capacitor area (prior art): 3mm x 5mm
Initial capacitance: 52 pF
Sensitivity: 0.133 pF/ kPa
 | Low threshold | High Threshold
Target force (g/N) | 4 / 0.0392 | 180 / 1.764
Target pressure (kPa) | 1.568 | 70.56
Capacitance change (pF/%) | 0.2085 / 0.4 | 9.384 / 18
CDC units | 910 | 40,999

## Slide 30: IRB PROPOSAL

## Slide 31: IRB proposal
Initial Draft X01:
  Draft sent to IRB Review Board on 10/13
Revision X02:
  Draft sent to IRB Review Board on 10/26
Revision X03:
  Draft sent to IRB Review Board on 10/29
Revision X04:
  Draft sent to IRB Review Board on 11/2

## Slide 32: Semmes Weinstein scale
4-180 grams of target force

## Slide 33: Block diagram-motherboard
SENSORY
SENSING
WEARABLE
Wired Digital Communication
NRF52832
FEEDBACK
ARMBAND
MOTHERBOARD
LIPO
BATTERY
BMS
INDUCTIVE
WIRELESS
CHARGING
SENSORY
SENSING CIRCUIT

## Slide 34: Block diagram-sensory sensing
SENSORY
SENSING
PRESSURE
SENSORS [3X]
CAPACITIVE
PRESSURE
SENSORS [3X]
CAPACTIVE
SENSING
CIRCUIT [3X]
FINGER STRAIN
SENSOR [3X]
RESISTANCE
STRAIN
SENSORS [4X]
RESISTIVE
SENSING
CIRCUIT [7X]
GRASP STRAIN
SENSOR [1X]
WIRED DIGITAL COMMUNICATION [4 LINES]
NRF52832
TEMPERATURE
SENSOR [3X]
RESISTANCE
TEMPERATURE
SENSORS [3X]
SENSOR GLOVE
MOTHERBOARD

## Slide 35: PCB design
~60mm x ~40mm

## Slide 36: Resistive sensing circuit
Up to 8 resistance sensors

## Slide 37: Capacitive sensing circuit
Up to 4 Capacitance sensors**

## Slide 38: Power circuit

## Slide 39: MCU circuit

## Slide 40: Wireless charging: Inductive
Hardware
  Old wireless charging coil:
    TDK 45-15F5-G (48.2mm x 32.2mm x 1.0mm)
  Current inductive coil:
    WR222230-26M8-G (22 mm x 22mm x 0.87 mm)
TRANSMITTER
RECEIVER COIL
RECEIVER
CIRCUITRY
FOOTPRINT:
15mm x 15mm

## Slide 41: Wireless charging: RF
RF Harvesting
  Battery and battery-less electronics
  7 mA continuous @ 0.5 meter
Companies contacted for development kit:
  Humavox
    Still waiting for a response
  Energous
    In conversation to receive 10 units of their receivers
  Powercast
    Digikey [PURCHASED]
      Transmitter (TX91503 $99.00)
        Battery-less receiver module (P2110B $48.33)
        Battery charging receiver module (P1110B $32.00)
        Development kit (P1110-EVAL-PS $420.00)

## Slide 42: HOUSEKEEPING

## Slide 43: Housekeeping
Electronics inventory:
  Google spreadsheet listing what is in stock at the lab
  Resistors, capacitors, inductors, ICs, etc…
  Reduce wait time to purchase parts

## Slide 44: Electronics development: schematic
Schematic generation
  nRF52 with supporting circuitry (2 hours)
  BLE antenna and impedance matching circuit (2 hours)
  Wireless charging (5 hours)
  Lithium battery charging (2 hours)
  UI interaction (1 hour)
  USB charging (2 hours)
  Biosensors: temperature, IMU, ECG (4 hours)
  Test points, connectors, jumpers (2 hours)
Estimated time: 20 hours
Actual time spent: 22 hours
*Not in scope for current phase of development.

## Slide 45: Progress to date: 6/28/2020

## Slide 46: Progress to date: 6/28/2020

## Slide 47: Progress to date: 6/28/2020

## Slide 48: Progress to date: 6/28/2020

## Slide 49: Electronics development: PCB design
PCB design
  Major component layout (2 hours)
  BLE antenna and impedance matching circuit (3 hours)
  Wireless charging (2 hours)
  Lithium battery charging (1 hours)
  Overall board construction(1 hour)
  USB charging (2 hours)
  Biosensors: temperature, IMU, ECG (3 hours)
  Test points, connectors, jumpers (1 hours)
  Gerber file generation (2 hours)
  BOM generation (2 hours)
Estimated time: 16 hours
Actual time spent: 20 hours

## Slide 50: Progress to date: 6/28/2020

## Slide 51: Progress to date: 6/28/2020

## Slide 52: Progress to date: 6/28/2020

## Slide 53: PCB manufacturing specifications
Dimensions: 99.568 x 76.454 mm
Layers: 2
Minimum trace/space: 6 mil
Minimum hole size: 12 mil
Thickness: 1.6mm
Solder mask: Green
Surface finish: HASL with lead
Copper weight: 1 oz
Material type: FR4

## Slide 54: PCA quote
PCB Manufacturing:
  Price: $20
  Manufacturer: PCBMinions
  Lead time: 4 days
BOM:
  Price: ~$289.63
  Board Quantity: 3
  Supplier: Digikey/Mouser
  Lead time: 1-5 days
Stencil:
  Price: ~46.25
  Supplier: OSH Stencil
  Lead time: 5 days

## Slide 55: Firmware development
Firmware
  Project and BLE stack (5 hours)
  Sensor drivers
    I2C driver (4 hours)
    Temperature driver (2 hours)
    ECG driver (6 hours)
    IMU driver (3 hours)
  USB & RTT data logging (5 hours)
    RTT data logging (2 hours)
    USB data logging (3 hours)
  DAQ slave protocol (1 hour)

## Slide 56: PROJECT MANAGEMENT

## Slide 57: Introduction
Purpose:
  Design a development board to test the functionality of different ‘plug-in-play’ sensors, wireless charging, and BLE communication. In the future, these sensors will be used in different combinations to build wearable devices.

## Slide 58: Project management
Project status updates will be emailed to YSK and Dr. Yeo on a weekly basis by Monday morning (ET)
A rolling PowerPoint presentation will be used to document results and act as a final report
  Each progress update will be timestamped, and the revision number will be incremented
Primary form of communication: email
  Zoom/phone calls can not be conducted from 12 PM – 8 PM ET, Monday – Friday (due to my work schedule)
Project documents and resource will be available through my Github repository
Time commitment: average 10 hours/week

## Slide 59: Project requirements
Requirements:
  Design a development board PCB to demonstrate capabilities to connect to different bio-sensors
  Development board must include the following functionality:
    Wireless and USB charging
    Battery management
    USB data communication and logging
    Use the nRF52832 MCU/SoC
    BLE connectivity and BLE GATT database
    Temperature sensing
    Motion sensing
    ECG
    Capability to add more sensors in the future

## Slide 60: Schedule overview
Electronics development
  Date range: 6/1–7/31
  Estimated hours to produce hardware: 49 hours
  Estimated hours to test/debug/qualify hardware: TBD
Firmware development
  Date range: 6/28-7/31
  Estimated hours to produce firmware: 33 hours
  Estimated hours to test/debug/qualify firmware: TBD
Any questions about the scope?
Let me know!

## Slide 61: PROJECT PROPOSAL

## Slide 62: Firmware development
Firmware
  Project and BLE stack (5 hours)
  Sensor drivers
    I2C driver (4 hours)
    Temperature driver (2 hours)
    ECG driver (6 hours)
    IMU driver (3 hours)
  USB & RTT data logging (5 hours)
Future development*
  BLE GATT database and data transfer (25 hours)*
  Sensors: fuel gauge, H.R./Sp02 (8 hours)*
  Data storage: EEPROM, Flash (8 hours)*
*Not in scope for current phase of development.

## Slide 63: Hardware testing
Firmware testing
  Firmware sensor functionality testing (8 hours)
Hardware testing
  Overall hardware testing (5 hours)
  Hardware sensor functionality testing (8 hours)
Future development*
  Power consumption*
  BLE range and connection testing*
*Not in scope for current phase of development.

## Slide 64: Block diagram
PYTHON DATA LOGGER
BLE TX/RX
IMU
TEMPERATURE
MCU/SoC
SERIAL TO USB CONVERTER
ECG
HR/Sp02**
PHERIPHERAL SENSORS
BATTERY GAUGE**
SERIAL
WIRELESS
CHARGING COIL
WIRELESS CHARGING CIRCUITRY
BMS
LIPO BATTERY
SERIAL
USB
POWER
POWER
POWER
POWER
RF
**Not in scope for current phase of development.
ANDROID PHONE**
BLE
OUT OF SCOPE

## Slide 65: PROJECT PROPOSAL: POWER CONSUMPTION TESTING
OUT OF SCOPE

## Slide 66: Power consumption testing [TBD]
Objective: Measure the power consumption of various sensors during operation, idle, and sleep mode
Log battery current and voltage at a defined sampling rate
Produce a power consumption model to determine battery life
OUT OF SCOPE

## Slide 67: Power consumption testing [TBD]
Nordic power profiler
  Nano-amp resolution, fast sampling rate
  See spikes of BLE transmission
OUT OF SCOPE

## Slide 68: Schedule flowchart: hardware
PROJECT
PROPOSAL
SCHEMATIC
GENERATION
PCB COMPONENT
LAYOUT
PCB ROUTING
& MFG. FILE GENERATION
FULL TURNKEY
PCB MFG.
HARDWARE
TESTING*
POWER CONSUMPTION TESTING*2
SENSOR FUNCTIONALITY TESTING*1
6/1–6/7; ~5 Hours
6/7–6/14; ~20 Hours
6/14–6/17; ~3 Hours
7/12–7/19; 5 Hours
7/19–7/31; 8 Hours
6/28–7/12; ~4 Hours
6/17–6/28; ~12 Hours
8/23–9/6; TBD
*Physical hardware needed to debug and perform testing.
1Hardware functionality test of the sensors will be performed. Performance and accuracy will not be qualified. 2Not in scope for current phase of development.
BLE RANGE & CONNECTION
TESTING*2
8/23–9/06; TBD
OUT OF SCOPE

## Slide 69: Schedule flowchart: firmware
PROJECT
PROPOSAL
PROJECT STACK
ORGANIZATION
SENSOR
DRIVERS
SERIAL & RTT DATA LOGGING W/ PYTHON
SENSOR FUNCTIONALITY TESTING*1
BLE GATT
DATABASE2
POWER
CONSUMPTION
TESTING*1,2
7/1–6/7; ~5 Hours
6/28–7/2; ~5 Hours
7/2–7/12; ~15 Hours
8/16–8/23; ~25 Hours
8/23–9/6; TBD
7/19–7/31; ~8 Hours
7/5–7/12; ~5 Hours
*Physical hardware needed to debug and perform testing.
1Firmware functionality test of the sensors will be performed. Performance and accuracy will not be qualified.
2Not in scope for current phase of development.
OUT OF SCOPE

## Slide 70: Component selection: MCU/BLE
MCU/BLE: nRF52832
  6 pin, 2.00 mm x 2.00 mm, WSON-6
  ±0.1°C from -20°C to +50°C
  2.4 GHz BLE
  Power = 4 dBm
  Sensitivity = -96 dBm
  RAM = 64 kB, flash = 512 kB
  Serial interfaces: I2C, UART, SPI
  Supply voltage: 1.7V ~ 3.6V
  $5.46 [X1], 1490-1067-1-ND, LINK
nRF52832

## Slide 71: Component selection: clocks
HF clock: US3200005Z
  4 pin, 1.60 mm x 1.20 mm, 4-SMD
  32.00 MHz, 8 pF, ±10 ppm
  $1.45 [X1], US3200005ZCT-ND, LINK
LF clock: ECX-1210
  1210 package, 1.20 mm x 1.00 mm, 2-SMD
  32.7680 kHz, 9 pF, ±20 ppm
  $1.52 [X1], XC2283CT-ND, LINK
ECX-1210
US3200005Z

## Slide 72: Component selection: USB/Serial
USB connector: 62910416121
  4 pin, 16.00 mm x 10.62 mm, SMT type B
  USB TYPE-B, USB 2.0, right angle
  $3.83 [X1], 732-3152-1-ND, LINK
USB/I2C converter: FTDI FT201XQ-R
  16 pin, 4.0 mm x 4.0 mm, 16-QFN
  USB to I2C converter
  3.4 Mbit/s
  Low power consumption (8 mA)
  $2.04 [X1], 768-1131-1-ND, LINK
FT201XQ-R
62910416121

## Slide 73: Component selection: Bio-sensors
Temperature: TI TMP117
  6 pin, 2.00 mm x 2.00 mm, WSON-6
  ±0.1°C from -20°C to +50°C
  16-bit resolution: 0.0078 °C, I2C
  $8.02 [X1], 296-51834-1-ND, LINK
IMU: TI BQ27621-G1
  14 pin, 2.5 mm x 3.0 mm, 14-LGA
  I2C communication
  9-axis, 16-bit resolution
  Full operation mode = 925 uA
  $4.98 [X1], 828-1057-1-ND, LINK
BMI160
TMP117

## Slide 74: Component selection: Bio-sensors
H.R. & Sp02: MAX30102
  14 pin, 5.6 mm x 3.3 mm, 14-OESIP
  Pulse oximetry and heart-rate
  Temperature compensation, I2C
  $9.51 [X1], MAX30102EFD+-ND, LINK
ECG: MAX30003
  28 pin, 5.0 mm x 5.0 mm, 28-TQFN
  SPI communication
  15.5-bit resolution
  Ultra-low power, single-channel integrated
  $8.21 [X1], MAX30003CTI+-ND, LINK
MAX30003
MAX30102
OUT OF SCOPE

## Slide 75: Component selection: wireless charging
Wireless charging IC: TI BQ51003
  28 pin, 1.90 mm x 3.00 mm
  QI v1.2 communication protocol
  93% AC-DC efficiency
  $3.00 [X1], 296-39952-1-ND, LINK
LiPo battery charger IC: TI BQ25100
  6 pin, 1.6 mm x 0.90 mm, 4.2V 1 cell
  1% Charge voltage accuracy
  10 mA to 250 mA charging current
  $2.24 [X1], 296-38373-1-ND, LINK
BQ25100
BQ51003

## Slide 76: Component selection: power circuit
Reed Switch: ABLIC S-5712A
  4 pin, 1.90 mm x 3.00 mm, SNT-4A
  Non-contact ‘power switch’
  High accuracy hall effect switch IC
  $1.07 [X1], 1662-1617-1-ND, LINK
Linear Regulator: Analog Devices ADP172
  4 pin, 0.95 mm x 1.03 mm, 4-WLCSP
  0.1V dropout voltage
  300 mA max current
  $1.09 [X1], ADP172ACBZ-3.0-R7CT-ND, LINK
ADP172
S-5712A

## Slide 77: Component selection: power circuit
Charging coil: Wurth 760308101219
  1 coil, 1 layer, D=15 mm, H=0.6 mm
  11.8 uH, 750 mΩ, t = ±10%
  $7.43 [X1], 732-9675-ND, LINK
Fuel gauge: TI BQ27621-G1
  9 pin, 1.62 mm x 1.58 mm, 9-DSBGA, 0.5 mm pitch
  I2C communication
  Dynamic voltage correlation algorithm
  Battery capacity, state-of-charge, voltage
  $2.21 [X1], 296-37460-1-ND, LINK
BQ27621-G1
760308101219
OUT OF SCOPE

## Slide 78: Component selection: TBD
Components still need to be selected:
  BLE antenna
  BLE impedance matching circuit/filter
  DCDC circuit

## Slide 79: Possible functionality
Possible functionality:
  Local data storage? (SD Card, EEPROM, etc…)
  Type of ECG connector? (Digikey part number…)
  NFC?
  Any other ideas?